Tel:  +44 (0)131 668 8411

Fax:  +44 (0)131 668 8412

Email:  vista@roe.ac.uk,

http://www.roe.ac.uk/atc/projects/vista/

##### Document Title: 	VISTA Verification Test Procedures

Document Number: 	VIS-PLARO-ATC-95000-0007

Issue: 	1.9

Date: 	29 Oct 2008

Sign-off Table on Following Page.

The information contained in this document is strictly confidential and is intended for the addressee only.  The unauthorised use, disclosure, copying, alteration or distribution of this document is strictly prohibited and may be unlawful.

Sign-Off Table

| Document Prepared By: | Brian Stobie Control Engineer | Signature and Date: |  |
| --- | --- | --- | --- |
| Document Approved By: | Andy Born Systems Engineer | Signature and Date: |  |
| Document Released By: | Alistair McPherson Project Manager | Signature and Date: |  |
| Accepted on behalf of QMUL : | Jim Emerson Principal Investigator | Signature and Date: |  |
| Accepted on behalf of ESO : | Martin Cullum | Signature and Date: |  |

Change Record

| Issue | Date | Section(s) Affected | Description of Change/Change Request Reference/Remarks |
| --- | --- | --- | --- |
| 1.0 | 02/03/07 | All | First Formal Issue. |
| 1.3 | 25/03/08 |  | Tidy up test comments. |
| 1.4 | 03/07/08 | All | Update by Andy Born: Additional detail in AJBorn-owned testing procedures Review of WJS-/GD-owned procedures to assess detail. |
| 1.5 | 11/07/08 | All | “First-pass” development by AJBorn of GBD/WJS procedures to turn into step-by-step procedures. |
| 1.6 | 12/07/08 | All | Incorporates updated GBD procedures, bullet-pointed by AJBorn. |
| 1.7 | 14/07/08 | All | Updates to WJS procedures. |
| 1.71 | 20/07/08 | 4.12.6, 5.6.5.8(1), 5.6.5.8(3) | Update to procedure for 5.6.5.8(1) with details of Flat Field illumination alternations; xrefs added to 5.6.5.8(3) and 4.12.6. |
| 1.72 | 21/07/08 | 5.6.5.8(1) | Added details of light level adjustments (still too high, but captured as a baseline). |
| 1.73 | 24/07/08 | 4.8.3 | Updated with details of procedure performed by GBD on 10-Jul-2008. |
| 1.74 | 25/07/08 | 4.12.2 | Updated with details of procedure performed by GDB on 25-Jul-2008. |
| 1.75 | 19/09/08 | All | Clean up fonts, etc |
| 1.8 | 06/10/08 | 4 | Add ‘Test result boxes to those without. Modify 5.1.10.8.4 for safety reasons as indicated. Add note to 5.110.8.2. Test 5.1.11.2(4) is superfluous. Add note to 5.1.17.10. Add ALT/AZ brake test, plus caveat. |
| 1.9 | 29/10/08 | All | Remove test results boxes to comply with discussions with Martin Cullum on 29/10/08. Re-write 5.1.11.4 to test Axial scale factors. |

Table of Contents

1	Introduction	6

2	Acronyms and Abbreviations	6

3	Applicable and Referenced Documents	6

4	VISTA Verification Test Procedures	7

4.1	------TVP / 4.8.3------IR Camera Field	7

4.2	------TVP / 4.9.1.2 ------System Image Quality	10

4.3	------TVP / 4.10.1.4------Pincushion Distortion	12

4.4	------TVP  / 4.12.2------Throughput	13

4.5	------TVP / 4.12.3------ Scattered Light	15

4.6	------TVP / 4.12.4(1)------Emissivity: Telescope + sky	17

4.7	------TVP / 4.12.4(2)------Emissivity : Telescope + Camera	19

4.8	------TVP / 4.12.5----Ghosting	20

4.9	------TVP / 4.12.6------ Light leakage	21

4.10	------TVP / 4.12.7------System Noise Characteristics	22

4.11	------TVP /  4.13.1.2------Telescope pointing	23

4.12	------TVP /  4.13.1.3(1)------Telescope offset pointing < 2°	24

4.13	------TVP / 4.13.1.3(2)------Telescope offset pointing (star remains on autoguider)	25

4.14	------TVP /  4.13.1.4------Re-Acquisition	26

4.15	------TVP / 4.13.2------Acquisition Time	27

4.16	------TVP /  4.14.1(1) and (2)------Open loop tracking	28

4.17	------TVP / 4.14.1(3)------Non-sidereal tracking	29

4.18	------TVP / 4.14.2(2)------Closed-Loop Tracking & Autoguider Ready Time	30

4.19	------TVP / 4.15.1----Exposure Length	31

4.20	------TVP /  4.15.2(1,2)------Exposure accuracy	32

4.21	------TVP / 4.15.3----Time Stamping	33

4.22	------TVP /  4.15.4.2------IR Camera – Re-start exposures	34

4.23	------TVP /  4.15.4.3------IR Camera – Multiple Integrations	35

4.24	------TVP /  4.15.4.4(1)------Rapid Sequence of IR Integrations	36

4.25	------TVP / 4.16.1.2------Stored data forms	37

4.26	------TVP / 4.16.2------Concurrent operation (disk storage)	38

4.27	------TVP / 4.16.3------ Archiving	39

4.28	------TVP / 4.16.5------ Data Storage	40

4.29	------TVP / 4.16.8(1)&(2)------ Local Data Reduction: Performance and Impact	41

4.30	------TVP /  4.16.8(3)------Local Data Reduction - Software Updates	42

4.31	------TVP /  4.16.8(4)------ Local Data Reduction - FTP Export	43

4.32	------TVP /  4.17(1)------ Thermal Control – Above Primary Mirror	44

4.33	------TVP / 4.17(2)------ Thermal Control – Below Primary Mirror	46

4.34	------TVP / 5.1.9.8.2 (2)------ Balancing / Residual balancing torque	48

4.35	------TVP / 5.1.10.8.1------ Kinematic Ranges	49

4.36	------TVP / 5.1.10.8.2------(Azimuth and Altitude) Operational Conditions	50

4.37	------TVP / 5.1.10.8.4------ Telescope Limits	52

4.38	------TVP / 5.1.10.10----ALT/AZ Brakes	54

4.39	------TVP / 5.1.10.11 (2)------ Altitude Auxiliary Drive / Useable with 1000 Nm OOB	55

4.40	------TVP / 5.1.11.2(4)------ Verify mirror figure can be modified	57

4.41	------TVP / 5.1.11.4------ M1 / Axial Actuator Scale Factors	58

4.42	------TVP / 5.1.11.6----M1 Restraint	59

4.43	------TVP / 5.1.17.5------ Slew Cass while tracking open-loop	60

4.44	------TVP / 5.1.17.6------(Cassegrain) Performance	61

4.45	------TVP / 5.1.17.10----Cass Brakes	62

4.46	------TVP / 5.1.17.12----Safety Locking	63

4.47	------TVP / 5.1.18.2.1(1)------ Guiding - User Selectable Frame Rate	64

4.48	------TVP / 5.1.18.2.1(3)------ Autoguider star probability	65

4.49	------TVP / 5.1.18.2.2------ Low Order Wavefront Sensor	67

4.50	------TVP / 5.3.3.3------ Filter Deployment	69

4.51	------TVP / 5.3.6.1(5)------ Readout noise	70

4.52	------TVP / 5.5.4.8(1)------ Hardware and Data Rates	72

4.53	------TVP / 5.6.4.3------- Emergency Lighting	73

4.54	------TVP / 5.6.4.4------ Protection Against Fire	74

4.55	------TVP / 5.6.4.5------ Enclosure Video and Audio Monitoring	75

4.56	------TVP / 5.6.5.4(4)------ Dome Rotation System – Stops Within 5s	76

4.57	------TVP / 5.6.5.6------ Windscreen – correct following of Telescope motion	77

4.58	------TVP / 5.6.5.7(1)------ Moonscreen – correct following of Telescope motion	78

4.59	------TVP / 5.6.5.8(1)------ Flat-field lighting – control from VLT Control Room	80

4.60	------TVP / 5.6.5.8(3)----Flat Field Illumination (Repeatable)	82

4.61	------TVP / 5.6.5.9(2)------ Dome Crane – Can Reach M1 Handling Location	84

4.62	------TVP / 5.6.5.9(11)------ Dome Crane ‘Out of Park’ switch	85

4.63	------TVP / 5.6.5.9(12)------ Dome Crane inhibit	86

4.64	------TVP / 6.1.1------ Control of Equipment from Control Building	87

4.65	------TVP / 6.1.6(2)------ Missing Data or Poor Quality Indicated in the Headers	88

4.66	------TVP / 6.1.6(3)------ All Faults Shall Be Logged with Relevant Details	89

4.67	------TVP / 6.1.8------ Noise interference from telescope mechanisms	90

4.68	------TVP / 7.3.5.3------Camera Installation and Removal	91

4.69	------TVP / 7.3.5.4------ Camera Intervention	92

4.70	------TVP / 7.4.12------ Inter-Subsystem Interlock Test	93

4.71	------TVP / Supplementary (1)------ Black Paint Adherence	95

4.72	------TVP / Supplementary (2)------ LOWFS Image Quality Check	96

4.73	------TVP / Supplementary (3)------ System Image Quality Check	97

# Introduction

This document contains the individual test procedures necessary to verify the requirements of the VISTA Technical Specification AD01 and in accordance with the VISTA Verification Plan AD02. Some of the tests are intended to be performed before a customer witness (ESO and/or QMUL) during the formal Verification Phase and others are performed during the Integration and Commissioning phases and may or may not be formally witnessed according to customer preferences.

# Acronyms and Abbreviations

AO		Active Optics

BOB		Broker for Observation Blocks

CASU		Cambridge Astronomical Survey Unit

DIMM		Differential Image Motion Monitor

DSS		Digital Sky Survey

FWHM	Full Width Half Maximum

GUI		Graphical User Interface

HOWFS	High-Order Wave-Front System

IPL		Instrument Preparation Laboratory

IR		Infra-Red

IRACE	Infra-Red Array Control Electronics

LOWFS	Low-Order Wave-Front System

LVDT		Linear Variable Differential Transformer

N/A		Not Applicable

RF		Radio-Frequency

SIQ		System Image Quality

TBC		To Be Confirmed

TBD		To Be Defined

TCS		Telescope Control System

TVP		Technical Verification Plan

VPO 		VISTA Project Organisation

# Applicable and Referenced Documents

| No | Title | Number & Issue |
| --- | --- | --- |
| AD01 | VISTA Technical Verification Specification | VIS-SPE-ATC-00000-0008 |
| AD02 | VISTA Technical Verification Plan | VIS-PLA-ATC-95000-0006 |

# VISTA Verification Test Procedures

## ------TVP / 4.8.3------IR Camera Field

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | IR Camera Field |
| Location of Test | Site |
| Applicable Documents: | VIS-SPE-ATC-06000-0004 |
| Preceding Verification Activities | N/A |
| Estimate of test duration | 1 hour + analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator + Scientist |

Test Equipment Required: Telescope , IR Camera, Data analysis workstation Test Procedure or Procedure reference: Test Equipment Required: Telescope , IR Camera, Data analysis workstation Test Procedure or Procedure reference: 1. Take images of a set of suitable pre-selected fields of known astrometric reference stars (2MASS coordinates are easily good enough for this purpose. 2. Stacking of the separate images to produce a deep sky flatfield image will provide a direct measurement of the radial variation of the flat-field, and hence a measurement of the Vignetting (1a). 3. Identify coordinates of pairs of stars near the center and edge of the field. Differencing the coordinates will yield a measurement of the plate scale and pixel scales at center and edge locations. 4. Differencing the coordinates of stars located at opposite corners of the FPA will yield a measurement of the field of view covered by the detectors. ================================================================== TEST PERFORMED 10-JUL-2008: 10/7/2008: Image centered on 18h00,-35d00 (VIRCAM_IMG_OBS193_0025.fits). The following simple measurements from the image allow a determination of the basic focal plane parameters. Display image as multi-fits file in ds9. Locate the extreme corners of chips 1 and 16 using the WCS information from the image headers, and identify a star close to each corner: Chip 1: RA=18:03:08.83, DEC=-35:29:44.44 at pixel (112,57) Chip 16: RA=17:56:53.89, DEC=-34:29:54.81 Image rotator at 180 degrees. Download 2MASS images of these two chips based on the reported positions of the chip centres, identify the two stars in these images and note the coordinates from ds9 - WCS reporting agrees between the 2MASS and VISTA images. Difference in Declination: 3610.4" Difference in Right Ascension: 375 seconds = 375x15xcos(35) = 4607" Length of IR focal plane diagonal = 5853" = 1.626 degrees. Second star on Chip1 at pixel (2000,61): RA=18:03:08.654, DEC=-35:19:05.96 Separation in declination = 10m38.48s = 638.48", pixel scale = 0.338 arcsec/pixel; plate scale = 59.14 microns/arcsecond. Upper corner of chip 6: star at pixel (2036,2029) has RA=18:00:26.175, DEC=-35:2:32.66 Second star at (104,2017) has RA=18:00:26.552, DEC=-35:13:32.42 Separation in declination = 10m 59.76s = 659.76", pixel scale = 0.34 arcseconds/pixel; plate scale = 58.57 microns/arcsecond. As second test is a measurement of the offset between the Cass. Rotator axis and the center of the IR focal plane. This was performed by observing several fields at 0 and 180 degree rotator angles, identifying objects on opposing detectors and comparing the pixel coordinates, with the following results: Chips at center of field: Detector 11, 180 degree rotation, star at Xdet=245, Ydet=67, Detector 6, 0 degree rotation, star at Xdet=1778, Ydet=1955 (would map to (270,93) on chip 11 if FPA center is perfectly located on the rotator axis. Xdiff=26 pixels, Ydiff=25 pixels Chips at edge of field: Detector 13: rotation 180 degrees, star at Xdet=424, Ydet=1928 Detector 4: rotation 0 degrees, star at Xdet=1593, Ydet=88 (would map to (1455,1960) on chip 13 id FPA center is perfectly located on the rotator axis. Xdiff=31 pixels, Ydiff=32 pixels. Within a measurement uncertainty of ~1 pixel, the location of the rotator center is then offset from its nominal location in the focal plane by Xoff=-14 pixels=280micron, Yoff=-14 pixels=280 micron. This corresponds to an offset of 3.3 arcseconds, compared to a specification limit of 8 arcseconds. Conclusion: The field of view, location, plate- and pixel-scales of the IR camera are within specifications. Author : GD

## ------TVP / 4.9.1.2 ------System Image Quality

| VISTA System | Telescope, IR Camera, Enclosure |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | System Image Quality |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | Pretty much everything: especially tracking, aO lookup tables, closed-loop aO, M1 supports all functional, thermal control and ventilation, DIMM operational. |
| Estimate of test duration | Setup: n/a . Test: 7 partial nights selected for good seeing, and varying wind and Moon phase. Analysis time: ~ 2 days. Reconfigure: n/a |
| Observing/Weather Conditions required: | nights of better than median seeing Low/Medium/High Windspeeds Varying Moon phase |
| Estimate of manpower required | Operator; Scientist Software engineer Systems engineer |

Test Equipment Required: Telescope , IR Camera, External DIMM on VISTA peak. Test Procedure or Procedure reference: This test requires: aO lookup tables in good state; all M1 supports operational; closed-loop LOWFS and autoguiding verified and operational; enclosure thermal control and vent doors operating normally; external DIMM operating on VISTA Peak. 1. “Baseline” SIQ Images: Take images with windspeed <= 10 m/s at the following configurations: i) Alt = 60 deg for Hour Angles between -1 to +1 hour ii) Alt at 80,70,60,50,40,30, 25 for Hour Angles between -1h to +1h . [NB: for a subset of these images, also take images with offsets 0.475 and 0.95 detector width in Y , and 0.95 and 1.9 detector width in X , for analysis of pincushion distortion, see 4.10.1.4 ] 2. “Fast Axis & Axis Reversal” SIQ Images: Take images at the following configurations: i) Alt = 87 and Az near South (for fast Az + Cass velocity); ii) Hour angle ~ -2.5 and +2.5, Dec ~ -30 (for slow Az / faster Alt). iii) HA ~ +3, Dec ~ -10 (for Cass reversal). 3. “Into Wind” and “Crosswind” rejection: Take images at the following configurations: i) Alt ~ 60 deg and pointing upwind in ~ 12 m/s wind. ii) cross-wind in ~ 12 m/s and ~ 16 m/s windspeed Reduction/Analysis Procedure: 1. Perform offline analysis of images to obtain FWHMs of stellar images across the IR focal plane. 2. Scale the DIMM results according to (lambda)^-0.2 (airmass)^0.6 to predict seeing at actual VISTA wavelength and airmass. 2. Subtract the scaled-DIMM results in quadrature from measured VISTA FWHM to obtain SIQ. Compare with specifications and analyse parameter dependencies. Software tools: IRAF may be used for measuring FWHM's and image profiles. Author : WJS

## ------TVP / 4.10.1.4------Pincushion Distortion

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Pincushion distortion |
| Location of Test | Site |
| Applicable Documents: | Tech Spec vXX |
| Preceding Verification Activities | As per SIQ tests. |
| Estimate of test duration | (parallel with SIQ) |
| Observing/Weather Conditions required: | As for SIQ testing |
| Estimate of manpower required | Operator. Scientist. Software engineer. |

Test Equipment Required: Test Procedure or Procedure reference: Utilise a representative sample of the SIQ fields of 4.9.1.2:, i.e. a set of IR Camera images in good seeing covering a representative range of other parameters. Also, for a few of the above fields with non-high windspeed, take some additional images shifted by 0.425, 0.95 and 1.9 detector widths in each of x/y to get extra cross-linkage between detectors. Reduction/Analysis Procedure: - Perform Astrometric fitting of stars using the CASU/IRAF/astrom/pipeline software to determine: i) the centre, and ii) and amplitude of the pincushion distortion pattern. - check for repeatability across fields. These images may also be used to verify pixel scale and field of view. Author : WJS

## ------TVP  / 4.12.2------Throughput

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Focal plane |
| Brief Title of Test | Throughput |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | Silver Coating of Mirrors |
| Estimate of test duration | 1 night + 1 Day analysis at least one photometric night |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator. Scientist . |

Test Equipment Required: IR Camera, telescope, photometric night, data reduction software. Test Procedure or Procedure reference: - Take observations of multiple fields of photometric standard stars (several per chip). - Repeat observations of each field spread over the range of zenith distance in steps of 80,70,60,50,40,30, 25deg ZD. Reduction/Analysis Procedure: 1. fit the observed instrumental flux for each star in each band to determine the extinction, zero points and colour terms for the system. 2. Determine the throughput of the system from the zero points. Engineering test performed by GBD 2008-07-25: Early test results from 1 J band observation of a 2MASS touchstone field: Single Star, 2MASS J=11.322 IRAF phot detects 109074 ADU above sky from a flatfielded image. Gain for chip 11 taken to be 4.62 from VIS-TRE-IOA-20000-0018 Effective wavelength for J band = 1.25 microns Effective bandwidth = 0.18 microns Magnitude zero point for J = -23.8 Effective collecting area of primary mirror seen by each pixel = 81602.8cm2 Predicted count from 11.322 magnitude star = 170,000, assuming 5% absorption in the atmosphere. Estimated throughput in J band = 64% Of the three primary science bands, J is 'expected' to have the lowest throughput as the detectors are usually dropping out there. This number implies that the detector q.e. is at least 90% in this band, and that everything else is as expect. Author : GD

## ------TVP / 4.12.3------ Scattered Light

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Scattered Light |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 dark photometric night (parallel with throughput) + 1 bright photometric night. (i.e. at least 2 weeks elapsed) |
| Observing/Weather Conditions required: | [Dark Night] + [Bright Night]: |
| Estimate of manpower required | Operator + Scientist |

Test Equipment Required: Data software and analysis software. Test Procedure or Procedure reference: [Dark Night]: - Observe a ‘blank’ field at high galactic latitude and as far from bright stars as can reasonably be achieved. - Take sufficiently deep single pawprint exposures of this field in each band that the sky background level can be obtained. - Observe a similar ‘blank field’, still at high latitude, but 1.5 degrees from a 1st mag star (need to find a star that is known to be 1st mag in JHK…) . Notes: - It should be possible to detect the residual sky gradient due to the presence of this source, and a comparison of the background level with the reference frames gives the result. - Sky background level can be determined as the average modal value of the counts in box samples over the field in regions free of bright objects. [Bright Night]: - Repeat the reference field from the Dark night when this field is well away from the moon. - Take a series of observations of blank fields in each filter at 5 degree intervals in moon-object separation, stepping up to 45 degrees. - Repeat the above comparison of the background level with the new reference field measurements. - Continue measurements of Bright Night up to 45 degrees away from the moon… Reduction/Analysis Procedure: - Pass data through CASU pipeline for flatfielding/dark subtraction. - Visual inspection of the processed images will be required to determine location of ghost images, although serious ghosts will show up in the image detection pass of the pipeline analysis. Author : GD

## ------TVP / 4.12.4(1)------Emissivity: Telescope + sky

| VISTA System | Telescope and IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Emissivity : Telescope + Sky |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 Night + 0.5 day analysis |
| Observing/Weather Conditions required: | a clear night when a VLT is taking Ks band imaging with either ISAAC or HAWK-I; |
| Estimate of manpower required | Operator. Scientist. |

Test Equipment Required: Test Procedure or Procedure reference: (NB: obtaining emissivity for a ‘perfectly baffled system’ is difficult; probably use VLT as a benchmark). - Select a night when a VLT is taking Ks band imaging with either ISAAC or HAWK-I; query in advance when this is planned to happen, and check with VLT telescope operator in real time. - Take near-simultaneous images with VISTA at similar airmass to VLT. Reduction/Analysis Procedure: - Analyse sky brightness in both VLT and VISTA images - apply a correction to VLT measurement to obtain a hypothetical value for “VLT with silver-coated mirrors” - Use 2MASS stars in VISTA field to define zeropoint; - from known pixel scale, deduce sky brightness in Ks mag/arcsec^2. Notes: Availability of data: any time VLT is imaging in Ks band , assuming we can 'borrow' the data without redistributing to get around proprietary rights issues (or just use VLT sky-brightness numbers from ESO QC pipeline). “Liability” issue: unlikely, since ESO presumably trust the VLT, and there are no practicable alternatives given the time-dependence of sky brightness. Author : WJS

## ------TVP / 4.12.4(2)------Emissivity : Telescope + Camera

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Emissivity : Telescope + Camera |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 0 night – in parallel with 4.12.4(2) 1 day analysis. |
| Observing/Weather Conditions required: | cloudless (photometric) night |
| Estimate of manpower required | Operator. Scientist. |

Test Equipment Required: Telescope and IR Camera Test Procedure or Procedure reference: - take a series of images in Ks filter (at medium to high galactic latitude) at a range of airmasses from 1.0 to 2.5, in steps of 0.2 airmass. - Repeat the above measurement range 3 times each, up-going and down-going. Reduction/Analysis Procedure: 1. Measure sky background for each image. 2. Perform linear fit of sky flux vs airmass to give ”telescope" flux (in photon/sec/arcsec^2) 3. Use 2MASS stars to convert to Ks mag/arcsec^2 and then to SI units using standard values for Vega. 4. Extrapolate to zero airmass to derive telescope emissivity. 5. Compare with computed value for black-body at ambient temperature. Notes: there is no obvious practicable experimental method to separate “Telescope” & “Camera” contributions, therefore for the individual contributions one must rely on analysis. Author : WJS

## ------TVP / 4.12.5----Ghosting

| VISTA System | Telescope and IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Ghosting |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 0 –> (1-2 nights, same dark night as 4.12.3.1) + 0.5d analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator + Scientist |

Test Equipment Required: Test Procedure or Procedure reference: - Identify a target field with at least one relatively bright (6th mag) star. - Make repeated observations of this field with the star dithered around the focal plane (probably only need to dither around a single quadrant, but NEED to include positions where the star is in the field, but NOT on a detector to cover reflections from the FPA plate…). - Run ‘SExtractor’ to find and remove bright images from all 16 detectors for each frame. - Search for large fluctuations in the background maps. - Subtract off the dark sky map from each image. - Removing all detected sources from the resultant image to give a sky image which can be thresholded to search for regions that exceed 50% of the nominal sky background (Persistence effects could contaminate this measurement, so adequate time intervals should be given between each star position.) - Repeat the dark sky measurement (from 4.12.3) – this should be done BEFORE moving the star around the field to avoid persistence effects Author : GD

## ------TVP / 4.12.6------ Light leakage

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Filter wheel |
| Brief Title of Test | Light leakage |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 0.5 day (daytime) (During Integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator. Scientist. |

Test Equipment Required: IR Camera on telescope, flat-field screen and lamps. NOTE: Refer to procedure for 5.6.5.8(1) for details of modified bulb installation and level adjustment. Test Procedure or Procedure reference: 1. Point telescope at flat-field screen 2. Turn lamps on 3. Take a set of 5-sec flat-field exposures in each broadband filter: i) Y ii) J, iii) H, iv) Ks. 4. Take a set of 300-sec “leak” images with lamps on and dark filter inserted. 5. Take a set of 300-sec “dark” images with lamps off. Reduction/Analysis Procedure: 1. Subtract “Dark” from “Leak” images to give “Leakage Flux”. 2. Sum the broadband images with suitable correction factor to account for extra bandwidth outside the filters to give an approximate equivalent of a “white light” exposure. 3. Divide the leakage flux by a scaled version of white light exposure to give a leak percentage. Author : WJS

## ------TVP / 4.12.7------System Noise Characteristics

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | System Noise Characteristics |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | (2 nights – 2 months) elapsed, parallel with SIQ + 0.5d Analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator + Scientist + VDFS |

Test Equipment Required: Telescope and IR Camera Test Procedure or Procedure reference: (*********** NOTE timescale of tests **************) - Select a field that will be observable at useful Zenith distances for a period of at least 2 months. - Observe 15 minute exposures of this field over a period of 3-4 nights to generate a 16h stacked exposure which should be processed through the standard data flow pipeline. - Stack the images incrementally to produce a curve of S/N vs time for a sample of objects on each chip. - Repeat these observations over a period of 2 months. - Construct multiple stacks of the reduced data frames to compare S/N vs time curves for the following subsets: - The original incremental 16h stack - An incremental 16h stack obtained at the end of the two month period. - An incremental 16h stack spaced across the two month period. Author : GD

## ------TVP /  4.13.1.2------Telescope pointing

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Telescope pointing |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | Lots |
| Estimate of test duration | 1 night + 1d Analysis + 1d Preparation |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software/system |

Test Equipment Required: None Test Script Not required - already in vtpom/pom. Test Procedure or Procedure reference: 1. Prepare a set of ~50 targets with guide stars. 2. Preset the telescope to each of these in turn and the guiding star positional error logged 3. Log the guide correction for each target guide star Reduction/Analysis Procedure: 1. Use TPOINT-based vtpom module to determine the blind pointing accuracy. (this is a standard ESO pointing test module) Success Criterion: < 3 arcsec rms Author : JMS

## ------TVP /  4.13.1.3(1)------Telescope offset pointing < 2°

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Telescope offset pointing < 2° |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | subsumed in absolute pointing test |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software/system |

Test Equipment Required: None Test Script Data analysis would be made more efficient using a new script. Test Procedure or Procedure reference: - Identify a set of targets; - For each target identify an offset target with guide star. - Preset to the initial target and log the guide star positional error - Preset the telescope to the "offset" target, and log the new guide star positional error - Repeat for each target. Reduction/Analysis Procedure: - The vector difference of these two guide errors is the offset pointing error. (Note: operational telescope offsets will not be performed; rather individual telescope presets will be performed because of the need to acquire new guide and active optics stars. There is currently no facility to "zero" the pointing error, before presetting to the nearby target.) Success Citerion: < 2 arcsec (95% confidence) Author : JMS

## ------TVP / 4.13.1.3(2)------Telescope offset pointing (star remains on autoguider)

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Telescope offset pointing (star remains on autoguider) |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 night +1d Analysis + 1d preparation |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software/system |

Test Equipment Required: None Test Script A new script would allow the procedure to be run efficiently. Test Procedure or Procedure reference: - Acquire a target such that one star appears on a science array and another on an autoguider. - Close the guiding loop and measure the position of the science array star - Move the telescope to a new target position but with the same guide star, and the guide loop closed. - Measure the new position of the science array star - Repeat for ~50 targets. Reduction/Analysis Procedure: - Compare the difference between these two positions in pixel units with that calculated from the offset. Success Citerion: < 0.5 pixels on science array (95% confidence) Author : JMS

## ------TVP /  4.13.1.4------Re-Acquisition

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Re-Acquisition |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software/system |

Test Equipment Required: None Test Script Test Procedure or Procedure reference: - Select a medium-dense star field at Dec ~ -40 , transiting near midnight. - Expose an image stack totalling 5 mins (to average out seeing) , recording the exact pixel coordinates of the guide star used. - Re-visit same field a few nights later at same hour angle (remember the 4 minutes/night UT-LST drift) - Re-use the same guide star and bring it back to the same pixel. Reduction/Analysis Procedure: - Compare the star positions on science arrays Success Criterion: Star positions on Science array should coincide; differences suggest variations in atmospheric refraction, AG-to-FPA differential flexure or Cass rotator zero-point . Author: WJS

## ------TVP / 4.13.2------Acquisition Time

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Axes |
| Brief Title of Test | Acquisition Time |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | 4.13.1.2 Absolute Pointing |
| Estimate of test duration | Setup: 0.5 hr Test: 1.0 hr Analysis: 0.5 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope plus Camera Test Procedure or Procedure reference: 1. Select a suitable sky target, and point the system so that the target is well positioned in the Guiding sensor. 2. Move the telescope off in Azimuth and Altitude so that the total angular position moved corresponds to the entries in the following table, and wait until the system has settled. 3. Start a timer (or read the timestamp information in the data off-line), bring the telescope back to re-position the target on to the same position on the guide sensor, and stop the timer when the target position error on the guider is less than 0.2 arcsec (i.e. the required 15s open-loop tracking accuracy). Slew Distance; Acquisition Time; Pass/Fail. 60 deg; ( < 45s ). 15 deg; ( < 30s ). 2 deg; ( < 10s ). 10 arcmin; ( < 6s ). 2 arcmin; ( < 3s ). 20 arcsec; (< 2s ). Author: BS

## ------TVP /  4.14.1(1) and (2)------Open loop tracking

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Open loop tracking |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 night + 0.5d analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software/system |

Test Equipment Required: None Test Script New script(s) for operation and analysis would make procedure more efficient. Test Procedure or Procedure reference: - Acquire a star on the autoguider but do NOT close the loop. - Log autoguider centroids at 10 Hz for a duration of 5 minutes. - Repeat the procedure at ~20 different positions, with different elevations and in/out of wind. Reduction/Analysis Procedure: - Plot centroids against time; - Calculate RMS centroid error, correct for seeing-induced image wobble and compare with Success Criteria below. Success Criteria: < 0.2 arcsec over 15s < 0.5 arcsec over 5 minutes. Author : JMS

## ------TVP / 4.14.1(3)------Non-sidereal tracking

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Non-sidereal tracking |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | In parallel with 4.14.1((1) and (2) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator Scientist |

Test Equipment Required: Test Procedure or Procedure reference: OPTION A: - Pre-select a near-earth asteroid moving at rate of 1 arcsec/s, - track it taking a series of short exposures Reduction/Analysis Procedure: Measure the centroids against time OPTION B: - Track low-density stellar field (e.g. galactic pole) at 1 arcsec/s non-sidereal rate. - Take a set of short exposures with DIT=2 sec, NDIT=5, separated by 10 sec pauses. Reduction/Analysis Procedure: - Star trails should appear as ‘lines’ with intensity dips along them. - Check the non-sidereal tracking by comparing exposures over the course of a few minutes. As long as a good WCS with plate scale and ZPN distortion is available, just reading off coordinates of trail end-points to the nearest pixel should be adequate Success Criterion: No specific criterion defined – implemented as a standard ESO TCS function. Author : WJS

## ------TVP / 4.14.2(2)------Closed-Loop Tracking & Autoguider Ready Time

| VISTA System | System –level – Tracking & Guiding |
| --- | --- |
| VISTA Sub-System | LOCSAG/Software/TCS/Telescope |
| Brief Title of Test | Closed-Loop Tracking & Autoguider Ready time |
| Location of Test | Site (on-sky) |
| Applicable Documents: |  |
| Preceding Verification Activities | - |
| Estimate of test duration | Setup: 0.25d (offline – selection of guide stars to use) Test: 1n Analysis time: 1d Reconfigure: none. |
| Observing/Weather Conditions required: | Good |
| Estimate of effort required | Systems Engineer Project Scientist Workpackage Manager - Software Telescope Operator |

Test Equipment Required: None. Test Procedure or Procedure reference: Define the set of stars to be used, from ESO online guide-star catalogue – ensure “realistic” distribution of magnitudes, elevations, positions wrt to galactic equator etc. For each guide-star in turn, select closed-loop guiding and measure the time taken to confirm closed-loop operation. Success critieria: - All candidate stars can be used for closed-loop guiding. - In each case, the time taken to close the loop is <1sec. Author : JMS

## ------TVP / 4.15.1----Exposure Length

| VISTA System |  |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test |  |
| Location of Test |  |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: *************Covered by tests in 4.15.2(1,2) following.************** Author : GD

## ------TVP /  4.15.2(1,2)------Exposure accuracy

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Detectors |
| Brief Title of Test | Exposure duration |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 Night + 0.5d analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator. Scientist. |

Test Equipment Required: Test Procedure or Procedure reference: OPTION A: (Method for longer exposures). - Point telescope at celestial equator, then stop it. - Take a series of exposures of varying DIT ~ 5s, 10s, 20s, 40s. Reduction/Analysis Procedure: - Bright stars will produce star trails at sidereal rate, 15.04 arcsec/sec; true exposure length can be calculated from length of trail and known plate scale. (Method for shorter exposures) - Take a series of images tracking a single star field with NDIT*DIT = 16*1s, 8*2s, 4*4s, 2*8s, 1*16s. Also 8*1s, 4*2s, 1*8s; 4*1s, 2*2s, 1*4s. Reduction/Analysis Procedure: 1. separate detector non-linearity from timing error by analysis of photometry (differential) – requires non-saturated stars of varying fluxes. Author : WJS

## ------TVP / 4.15.3----Time Stamping

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Exposure time stamp |
| Location of Test |  |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 hour. |
| Observing/Weather Conditions required: | Any. |
| Estimate of manpower required | Operator Scientist. |

Test Equipment Required: Test Procedure or Procedure reference: Start a 20 sec DIT in J-band on celestial equator. Approximately in mid-exposure, hit the STOP TRACK Button on the TCS screen; this event with UT should appear in TCS event logs. Measure star trail lengths on resulting partly-trailed image; using trail lengths and Exposure end-time in header can deduce UT of the STOP event. Compare the calculated and logged UTs ; this checks that IRACE and the TCS have a consistent clocks. Author : WJS

## ------TVP /  4.15.4.2------IR Camera – Re-start exposures

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Detectors |
| Brief Title of Test | IR Camera - Re-start Exposures |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: Tested at Garching by S.Beard (Average data storage overhead = 3.6 sec, email from SB Monday November 19. 2007 2:34 PM). ‘Should be repeated using new instrument workstation hardware at Paranal.’ Author: GD

## ------TVP /  4.15.4.3------IR Camera – Multiple Integrations

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Detectors |
| Brief Title of Test | IR Camera – Multiple Integrations |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: By Design/inspection, inherent property of IRACE system, Set DIT=10 and make integrations happen every 10s. Email from SB Monday November 19. 2007 2:34 PM Author: GD

## ------TVP /  4.15.4.4(1)------Rapid Sequence of IR Integrations

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Detectors |
| Brief Title of Test | IR camera – Rapid Sequence of IR Integrations |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: Refer to Email from SB Monday November 19. 2007 2:34 PM, do at same time as test 4.15.4.2. Result dependant on using ‘Double’ or ‘DoubleReadReset’ readout modes. Author: GD

## ------TVP / 4.16.1.2------Stored data forms

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Stored data forms |
| Location of Test | anywhere |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 10 minutes |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Scientist |

Test Equipment Required: Test Procedure or Procedure reference: - Use IMG_GEN_TEC.obd from BOB - change data type between Double and Uncorrelated and change the number of co-add - Take a range of data frames and verify that the resultant data are consistent with the requested readout modes. (Observation 10-Jul-2008: Stored data are indeed consistent with chosen readout modes.) The acquisition rate can be inferred by timing the readouts (this has already been done at Garching). The data capacity can be verified by typing ‘df’ on the instrument workstation. (Observation 10-Jul-2008: The data capacity of the workstation is now 5200 data frames. With a data rate of 1 image every 10 seconds for 14 hours, the maximum amount of data that can be observed in a single night is 5040 data frames.) Author : GD

## ------TVP / 4.16.2------Concurrent operation (disk storage)

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Concurrent operation (disk storage) |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 4 hours |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator / Software |

Test Equipment Required: Test Procedure or Procedure reference: A: This is a standard feature of the IRACE system. The behaviour can be verified by filling the disk (or by redirecting the data storage to a much smaller disk and filling that) and then trying to take an exposure. Test performed July 8th 2007 with expected behaviour. Success criterion: The Operator should be warned that the disk is full, before the exposure is commenced. B: This feature should now have been added to IRACE: - Schedule two OBs of different fields with different filters back-to-back. - Monitor the IRACE Engineering GUI screen to determine the point at which the exposure is complete, and the OS control GUI to determine the point at which the merged file is written to disk. Success criterion: The Telescope should have moved before the data-write is complete. Author : GD

## ------TVP / 4.16.3------ Archiving

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Archiving |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 day offline |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | ESO Data Flow Rep. /CASU |

Test Equipment Required: Test Procedure or Procedure reference: Process a full set of observations through the summit QC software or the VDFS software. Alternatively CASU have a version of FITSCHECK which will verify the header contents. Note: it should also be verified that the values written into the FITS headers are those which correspond to the current instrument/ telescope/enclosure configurations. This can be achieved by inspection of the headers. Success criterion: Headers are correct. Author : GD

## ------TVP / 4.16.5------ Data Storage

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Data Storage |
| Brief Title of Test | Data Storage |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 10min (During Integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Software |

Test Equipment Required: Test Procedure or Procedure reference: Verify that there is at least 10Tb of on-line storage available to the instrument workstation. Success criterion: Required storage available. Author : GD

## ------TVP / 4.16.8(1)&(2)------ Local Data Reduction: Performance and Impact

| VISTA System | IR Camera, Data Handling, Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Local Data Reduction: Performance and Impact |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 night (During Integration) |
| Estimate of manpower required | 1 Operator 1 Software |

Test Equipment Required: None Test Script Data analysis would be made more efficient using a new script. Test Procedure or Procedure reference: 1. Initiate such that data are acquired continuously at the full performance rate (one exposure every 10s). 2. Select a recent observation and analyse this on the offline workstation. Reduction/Analysis Procedure: Verify that this did not impact the observing process by analysing the observing log (timestamps of exposures). Success Criterion: IR Camera acquires data at rate of 1 exposure every 10s, whilst data are analysed on offline workstation. Author : JMS

## ------TVP /  4.16.8(3)------Local Data Reduction - Software Updates

| VISTA System | Data Handling |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Local Data Reduction - Software Updates |
| Location of Test |  |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 0 |
| Estimate of manpower required | NA |

Test Equipment Required: None Test Procedure or Procedure reference: This requirement will not be tested as it is essentially tautological. The requirement is that "authorised personnel" i.e. ESO staff shall be able to install software on what is actually ESO standard hardware, procured and configured by ESO. Success Criterion: N/A Author : JMS

## ------TVP /  4.16.8(4)------ Local Data Reduction - FTP Export

| VISTA System | Data Handling |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Local Data Reduction - FTP Export |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 hour (During Integration) |
| Estimate of manpower required | 1 Software |

Test Equipment Required: None Test Procedure or Procedure reference: 1. Select a data file on the Offline Workstation and transfer it to a computer outside the firewall by means of an FTP put operation (the external computer will need to run an FTP server). 2. Transfer a data file as an email attachment. Success Criterion: File can be read on the external computer. Author : JMS

## ------TVP /  4.17(1)------ Thermal Control – Above Primary Mirror

| VISTA System | Thermal Control |
| --- | --- |
| VISTA Sub-System | Enclosure, M2 Unit |
| Brief Title of Test | Thermal Control – Above Primary Mirror |
| Location of Test | Site |
| Applicable Documents: | None. |
| Preceding Verification Activities | None. |
| Estimate of test duration | (In parallel with 4.17(2)) Setup: 0.5d (stabilising time) Test: 1n (periodic temp measurement) Analysis time: n/a Reconfigure: n/a |
| Observing/Weather Conditions required: | Need a night with “lowish” windspeeds (Note: experimentation with Vent doors will be required to obtain required internal windspeed range). |
| Estimate of effort required | Systems Engineer Project Scientist Electrical Technician |

Test Equipment Required: Remote temperature measurement probe. Use Subsystem thermometer sensor readings or remote thermometer. Use Telescope Structure anemometers for consistency and convenience Test Procedure or Procedure reference: - Set the system into normal daytime “sleep” mode with all Enclosure thermal conditioning systems running as normal. - Wait for minimum of 6 hours to allow internal temperatures to stabilise. - Perform normal pre-sunset ventilation system operations to commence enclosure flushing. - Open observing slit doors - Rotate Enclosure and deploy windscreen, vent doors etc. to obtain ~2m/s windspeed within dome. - Commence observing sequence, with slewing in Alt/Az/Cass every ~15mins - At 1hr intervals, measure and log temperature of Telescope/M2 Unit/Enclosure structures/components above primary mirror level (pay particular attention to any sited in locations that may give rise to thermal plumes: Windscreen motor, Enclosure cabinets, brake resistors etc, Telescope Altitude motors, M2 Unit surfaces Reduction/Analysis Procedure: - Compare temperature deviation from ambient with the spec. Note: images gathered should be logged and archived, however those taken around the time of temperature measurement may be contaminated from the remote thermometer probe. Author : AJB

## ------TVP / 4.17(2)------ Thermal Control – Below Primary Mirror

| VISTA System | Thermal Control |
| --- | --- |
| VISTA Sub-System | Enclosure, M2 Unit |
| Brief Title of Test | Thermal Control – Below Primary Mirror |
| Location of Test | Site |
| Applicable Documents: | None. |
| Preceding Verification Activities | None. |
| Estimate of test duration | (In parallel with 4.17(1)) Setup: 0.5d (stabilising time) Test: 1n (periodic temp measurement) Analysis time: n/a Reconfigure: n/a |
| Observing/Weather Conditions required: | Need a night with “lowish” windspeeds. (Note: experimentation with Vent doors will be required to obtain required internal windspeed range). |
| Estimate of effort required | Systems Engineer Project Scientist Electrical Technician |

Test Equipment Required: Remote temperature measurement probe. Use subsystem thermometer sensor readings or remote thermometer. Use Telescope Structure anemometers for consistency and convenience Test Procedure or Procedure reference: - Set system into normal daytime “sleep” mode with all Enclosure thermal conditioning systems running as normal, - Wait at least 6 hours to allow temperatures to stabilise. - Perform normal pre-sunset ventilation system operations to commence enclosure flushing. - Open observing slit doors - rotate Enclosure and deploy windscreen, vent doors etc. to obtain ~1m/s windspeed within dome. - Commence observing sequence, with slewing in Alt/Az/Cass every ~15mins. - At 1hr intervals, measure temperature of Telescope/Enclosure structures/components below primary mirror level (pay particular attention to any sited in locations that may give rise to thermal plumes: M1 Cell, M1 Actuators, IR Camera rear surfaces. Reduction/Analysis Procedure: Compare temperature deviation from ambient with the spec. Note: images gathered should be logged and archived, however those taken around the time of temperature measurement may be contaminated from the remote thermometer probe. Author : AJB

## ------TVP / 5.1.9.8.2 (2)------ Balancing / Residual balancing torque

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Balancing / Residual balancing torque |
| Location of Test | Site |
| Applicable Documents: | Altitude Axis Balancing Procedure VIS-PRO-VER-01001-0691 |
| Preceding Verification Activities |  |
| Estimate of test duration | Setup: 0 hr Test: 2 hr Analysis time: 2 hrs (offline) Reconfigure: n/a |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Mech Tech |

Test Equipment Required: Telescope, integrated altitude axis LCU Test Procedure or Procedure reference: - Check the altitude axis balance as per the Altitude Axis Balancing Procedure. - Drive the altitude axis from zenith to horizon and back at constant velocity, recording the altitude motor current. Reduction/Analysis Procedure: Analyse variation of current over alt range to estimate the residual balancing torque. TEST DATE; PASS / FAIL. LOCATION OF TEST RESULTS. Author : PFJ

## ------TVP / 5.1.10.8.1------ Kinematic Ranges

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Axes |
| Brief Title of Test | AZ, ALT Kinematic Ranges |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | n/a |
| Estimate of test duration | (During integration) Setup: 0.5 hr Test: 1.0 hr Analysis: 2.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope and IR Camera Test Procedure or Procedure reference: (a) Verify control system zero points and axis definitions: (b) Command maximum angles (increasing software limits if necessary) and record reported angles. Author: BS

## ------TVP / 5.1.10.8.2------(Azimuth and Altitude) Operational Conditions

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Axes |
| Brief Title of Test | (Azimuth and Altitude) Operational Conditions |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | 5.1.10.8.1 Kinematic Ranges |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 2.0 hr Analysis: 2.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope and IR Camera Test Procedure or Procedure reference: Telescope Azimuth and Altitude drives driven under software control to verify functional velocity and accelerations (actual system tracking capability including errors is verified in Procedure 4.14.1): This requires logging of position vs. time and later analysis, in addition, maximum tracking velocity and accelerations cannot be verified by normal sky objects. Procedure TBD. NOTE – interface scaling and tolerances may result in a ‘Fail’ for Slew Velocity. Author :BS

## ------TVP / 5.1.10.8.4------ Telescope Limits

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Axes |
| Brief Title of Test | Telescope Limits |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | 5.1.10.8.1 Kinematic Ranges |
| Estimate of test duration | (During integration) Setup: 1.0 hr Test: 3.0 hr Analysis: 0 hr Reconfigure: 1.0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator 1 Safety Authority |

Test Equipment Required: Telescope. NOTE (1) : Tested during VRSI On-Site Acceptance Tests - Early ESO warning if witnessing required. Software Limits are tested as part of Axis LCU Site Integration activities. Test Procedure or Procedure reference: (VRSI) VISTA MCS Integrated Test Procedure, VIS-PRO-VER-01001-9008, S/4.4 The telescope software limits can be verified using the TCS under LCU control. The telescope drives can be driven using the Vendor maintenance system through their functional ranges to verify the ‘Vicinity’ limits. To test the further hardware ‘Pre-interlock’ limit would require adding wires inside the PDU electronics, which carries some risk. As all the limits are hardware cam-operated switches, and have already been verified at Telescope acceptance, this risk is not necessary. NOTE Each axis has 4 sequential hardware-switched limits – - Velocity limit, which limits velocity to 0.5 deg/s. - Vicinity limit, which automatically reverses the drive to back out of the limit. - Pre-interlock limit, which disables the motor power drives. - Interlock Limit, which activates breakers to remove motor electrical power and applies the (fail-safe) air brakes. Author: BS

## ------TVP / 5.1.10.10----ALT/AZ Brakes

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | ALT / AZ |
| Brief Title of Test | ALT/AZ Brakes |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 10 minutes. |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Elect Tech 1 Systems / Electronic Engineer |

Test Equipment Required: Telescope & Axis LCU’s or VRSI PMU Test Procedure or Procedure reference: NOTE : This test inevitably stresses both the electrical and mechanical systems of the telescope. Suggest ESO accepts tests already done during telescope acceptance. Otherwise, set the axis angle to the mid-range of travel. If using the LCU, request an angle 20 deg greater than starting angle. If using PMU, set velocity to positive maximum. When the axis reaches maximum velocity, apply an Emergency Stop. Verify brakes are applied, breakers trip, and axis stops. Author : BS

## ------TVP / 5.1.10.11 (2)------ Altitude Auxiliary Drive / Useable with 1000 Nm OOB

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Altitude Auxiliary Drive / Useable with 1000 Nm OOB |
| Location of Test | Site |
| Applicable Documents: | Altitude Axis Balancing Procedure VIS-PRO-VER-01001-0691 |
| Preceding Verification Activities |  |
| Estimate of test duration | Setup: 1.0 hr Test: 2.0 hr Analysis: 0 hr Reconfigure: 1/2 hr |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Mech Tech |

Test Equipment Required: Telescope, integrated altitude axis LCU, force gauge Test Procedure or Procedure reference: 1. Confirm the altitude axis is balanced as per the Altitude Axis Balancing Procedure 2. Park the telescope at zenith. 3. Station a Technician on the Alt platform to operate the auxiliary drive. 4. Station a second Technician to operate the Alt brake override button (“dead man’s handle”). 5. Remove the altitude auxiliary drive hand crank from storage position and engage the auxiliary drive. 6. When ready to commence the test, the second Technician engages the alt brake override button 7. The technician shall drive the telescope alt axis from zenith to horizon, and back to zenith. 8. Measure the required force on the hand crank to move the axis at a minimum of 3 points in the range of travel using a force gauge. (An RFW is anticipated as the max out of balance torque possible with the manual drive is <1000Nm.) Author : PFJ

## ------TVP / 5.1.11.2(4)------ Verify mirror figure can be modified

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | M1 support |
| Brief Title of Test | Verify mirror figure can be modified |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 0.5 hr Analysis: 2.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope and (Test) Camera Test Procedure or Procedure reference: This test is superfluous, as many examples of modified M1 image data are available from many nights of commissioning testing of VISTA. Author: BS

## ------TVP / 5.1.11.4------ M1 / Axial Actuator Scale Factors

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | M1 / Axial Actuator Scale factors |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 6 hours (During integration) |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator |

Test Equipment Required: Test Procedure or Procedure reference: Run M1 diagnostic test script TBD (Malcolm Stewart). Verify nominal scale factors of axial actuators by analysing results. Author : BS

## ------TVP / 5.1.11.6----M1 Restraint

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | M1 |
| Brief Title of Test | M1 Restraint |
| Location of Test | Site |
| Applicable Documents: | VIS-PRO-VER-01001-9008 Rev A |
| Preceding Verification Activities |  |
| Estimate of test duration | 2 hrs |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Elect. Tech. |

Test Equipment Required: Telescope, Integrated Altitude and M1 LCU Test Procedure or Procedure reference: Repeat VRSI test procedure : VISTA MCS, Integrated Test Procedure VIS-PRO-VER-01001-9008, 11/11/05 Para 5.10 Mirror restraining test. Author : PFJ

## ------TVP / 5.1.17.5------ Slew Cass while tracking open-loop

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Cass Axis |
| Brief Title of Test | Slew Cass while tracking open-loop |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | 5.1.10.8.1 Kinematic Ranges |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 0.5 hr Analysis: 0.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope and (Test) Camera Test Procedure or Procedure reference: (1) Send open-loop track demands to the Azimuth and Altitude axes. (2) Send a slew demand to the Cassegrain axis, and verify that the AZ and ALT axes continue to track without interruption. NOTE – Informally verified during commissioning, 11/08/0-8, see Trac Log #195 Author: BS

## ------TVP / 5.1.17.6------(Cassegrain) Performance

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Cass Axis |
| Brief Title of Test | (Cassegrain) Performance |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | 4.13.1.2 Absolute Pointing 5.1.10.8.1 Kinematic Ranges |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 1.0 hr Analysis: 1.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Telescope and (Test) Camera Test Procedure or Procedure reference: Repeat procedure 5.1.10.8.2 for the CASS axis. Author: BS

## ------TVP / 5.1.17.10----Cass Brakes

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Cassegrain |
| Brief Title of Test | Cass Brakes |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 10 minutes. |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Elect Tech 1 Systems / Electronic Engineer |

Test Equipment Required: Telescope & Axis LCU’s or VRSI PMU Test Procedure or Procedure reference: NOTE : This test inevitably stresses both the electrical and mechanical systems of the telescope. Suggest ESO accepts tests already done during telescope acceptance. Otherwise, set the axis angle to the mid-range of travel. If using the LCU, request an angle 20 deg greater than starting angle. If using PMU, set velocity to positive maximum. When the axis reaches maximum velocity, apply an Emergency Stop. Verify brakes are applied, breakers trip, and axis stops. Author : BS

## ------TVP / 5.1.17.12----Safety Locking

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Cassegrain |
| Brief Title of Test | Safety Locking |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 2 hrs |
| Observing/Weather Conditions required: | n/a |
| Estimate of manpower required | 1 Operator 1 Mech. Tech. |

Test Equipment Required: Telescope and axis LCU’s Test Procedure or Procedure reference: Position the axis using LCU control at the stow position. Activate the Stow Pin. Ensure that ‘stow pin engaged’ interlock is activated and ‘axis stowed’ signals are reported to the LCU. Author : PFJ

## ------TVP / 5.1.18.2.1(1)------ Guiding - User Selectable Frame Rate

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Guiding - User Selectable Frame Rate |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 2 hour |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software |

Test Equipment Required: None Test Procedure or Procedure reference: - Acquire a star on the autoguider - Close the guide loop and log the timestamped guide corrections. - Select the following guide rates: 10 Hz, 5 Hz, 2 Hz and 1 Hz. Reduction/Analysis Procedure: Examine the timestamps. Success Criterion: Timestamped guide corrections are at the rate demanded by the user. Author : JMS

## ------TVP / 5.1.18.2.1(3)------ Autoguider star probability

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Autoguider |
| Brief Title of Test | Autoguider star probability |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 Night + 1 day analysis |
| Observing/Weather Conditions required: | Better than median seeing. Full moon, no cirrus. |
| Estimate of manpower required | Operator, Scientist, Software engineer. |

Test Equipment Required: Test Procedure or Procedure reference: Empirically measure the faintest guide star magnitude permitting acceptable guiding performance in the conditions specified above, as follows: - Using a range of telescope pointings and/or different guide star(s) from candidates list, select guide-stars with I-band magnitudes ~ 14.0 to 17.0 in 0.25 mag steps approx. - For each guide star above, take a 3 minute exposure with autoguider running and logging centroid data, but guide loop not closed. - record rms offset of guide star centroid from mean position Reduction/Analysis Procedure: - Plot guide-star positional rms error vs guide-star magnitude - locate 'knee' where rms increases significantly due to autoguider photon and/or readout noise. Given the magnitude determined above, calculate guide-star probability assuming Poisson statistics as follows: Use default guide-star catalogue (GSC-II) and count stars matching above selection criteria in about 10 x 1-deg^2 patches near the South Galactic Pole at 00h50m, -27 deg. This gives guide star density; then, known AG useful area and Poisson statistics gives probability of usable guide star for random pointing near SGP. Author : WJS

## ------TVP / 5.1.18.2.2------ Low Order Wavefront Sensor

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Low Order Wavefront Sensor |
| Brief Title of Test | LOWFS Performance |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | LOWFS software. |
| Estimate of test duration | 2 Nights + 2d analysis |
| Observing/Weather Conditions required: | Good seeing. |
| Estimate of manpower required | Operator, Scientist , Software engineer. |

Test Equipment Required: Start from a known "good" active optics setup. Test Procedure or Procedure reference: TEST 1 – motion/aberration consistency: 1. Track 1h East of meridian at Alt ~ 60 degrees. 2. Take simultaneous HOWFS and LOWFS measurements with a range of exposure times 20s – 120s in 20s steps. - Deliberately offset M2 via Active Optics screen by applying single offsets of +/-250nm, +/- 500nm, and +/-1000nm in each term as follows, taking simultaneous HOWFS and LOWFS measurements with a range of exposure times 20s – 120s in 20s steps.: i) Z4 (focus) ii) Z5/Z6 (tilt/astig) iii) Z7/Z8 (coma); - Deliberately offset M2 via Active Optics screen by applying pair-wise offsets of +/-250nm, +/- 500nm, and +/-1000nm in each term as follows, taking simultaneous HOWFS and LOWFS measurements with a range of exposure times 20s – 120s in 20s steps.: i) Z4 (focus) ii) Z5/Z6 (tilt/astig) iii) Z7/Z8 (coma); Reduction/Analysis Procedure: - Compare results before and after each M2 movement. - Verify that the observed aberrations are consistent with introduced M2 motions. TEST 2 – star quality: 1. Pre-select a few “intentionally bad” LOWFS stars, e.g. binaries of similar brightness and separations ~ 2 – 5 arcsec. 2. Take observations at a suitable pointing such that the binary is the chosen LOWFS star [if necessary, edit .paf file to move bad star to the first choice] Reduction/Analysis Procedure: Confirm that the LOWFS analysis software reports bad data. TEST 3 – M2 motion: Deliberately offset M2 by a substantial offset between the end of one observation block and the start of the next. Reduction/Analysis Procedure: Confirm that an operator warning is issued after the next LOWFS cycle completes. Author : WJS

## ------TVP / 5.3.3.3------ Filter Deployment

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | Filter Wheel |
| Brief Title of Test | Filter Deployment |
| Location of Test | Anywhere |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 hour (During integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator or systems or Scientist |

Test Equipment Required: Stopwatch Test Procedure or Procedure reference: - Request an adjacent filter and time the response of the system. - Request a move from a narrow-band filter to an adjacent-but-one filter that would cross a broad-band filter. - Repeat to include clockwise and anti-clockwise movements. Reduction/Analysis Procedure: Verify that the time taken to move through 6 filter positions is <=60s. Author : GD

## ------TVP / 5.3.6.1(5)------ Readout noise

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System | IR Detectors |
| Brief Title of Test | Readout noise |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 4 hrs + 4hrs analysis |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Operator Scientist VDFS software expert. |

Test Equipment Required: Test Procedure or Procedure reference: 1. Take a set of dark frames daytime with telescope parked 2. Take a second set of dark frames at night with telescope and dome tracking. Reduction/Analysis Procedure: 1. Measure noise level on each set of frames 2. Compare to check any excess noise arising from telescope and/or dome. Test procedure (2): 1. Take several sets of exposures at a single telescope pointing, with a low-background filter (Z or NB118) 2. Repeat over a range of exposure times DIT, approx 10 sec to 300 sec. Reduction/Analysis Procedure: 1. Analyse noise level vs mean background level 2. Confirm consistency with predicted readout + sky noise. Note: Requires use of the CASU algorithms including “de-curtaining”. Author : WJS

## ------TVP / 5.5.4.8(1)------ Hardware and Data Rates

| VISTA System | Data Handling |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Hardware and Data Rates |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 4 hours (During integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Software |

Test Equipment Required: None Test Procedure or Procedure reference: Capacity Test: Execute the command df or equivalent on each workstation. Data Rate Test: - Initiate a series of observations lasting 30 minutes at the full rate, 1 exposure per 10s 2. Check the number and total size of the results data files using the command ls -l on each of the Instrument, Data Handling and Software workstations. Notes: (a) Perform a separate test to verifty the storage rate requirement of 54 MB/s on the Instrument workstation, and to store data within 5s. (b) The User Workstation's appearance in this list is incorrect. (c) ESO are responsible for the performance of the Archive system. Success Criteria: Disk capacities on each workstation are at least 2.8 TB. Data handling, and software workstations keep up with the incoming data. Author : JMS

## ------TVP / 5.6.4.3------- Emergency Lighting

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Emergency Lighting |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: COMPLETED at EIE Acceptance, VIS-PRO-EIE-01000-0232, VIS-LIS-EIE10100-0332

## ------TVP / 5.6.4.4------ Protection Against Fire

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Fire Alarm System |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Gas source for Smoke head testing Test Procedure or Procedure reference: (1)Test 20% of smoke heads using gas source : Verify all sounders go off. Verify HVAC is switched off. Verify status signal appears at VLT Fire System. (2) Test 20% of call points – remove break cover by dismantling panel top. Verify as for (1)

## ------TVP / 5.6.4.5------ Enclosure Video and Audio Monitoring

| VISTA System | Enclosure (and Networking) |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Enclosure Video and Audio Monitoring |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 2 hours (During integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 ANO |

Test Equipment Required: None Test Procedure or Procedure reference: - Position someone at appropriate locations in the enclosure to wave their arms and shout. - Position someone else in the Paranal Control Room to observe them on the monitor and loudspeaker. - Move the Telescope. Success Criteria: The people in the Enclosure and the moving telescope are heard, and seen to move. Author : JMS

## ------TVP / 5.6.5.4(4)------ Dome Rotation System – Stops Within 5s

COMPLETED

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Dome |
| Brief Title of Test | Dome rotation |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: Scanned document VIS-PRO-EIE-10100-0149 / p4.bmp From Livelink https://livelink.roe.ac.uk/Livelink/livelink.exe/properties/2152552

## ------TVP / 5.6.5.6------ Windscreen – correct following of Telescope motion

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Windscreen, ECS. |
| Brief Title of Test | Windscreen – correct following of Telescope motion |
| Location of Test | Site |
| Applicable Documents: | None. |
| Preceding Verification Activities | Testing of manufacture, operation and range of motion is covered under Enclosure acceptance testing. |
| Estimate of test duration | (During integration) Setup: 1hr (daytime) Test: 1hr (daytime, OSD closed) Analysis time: n/a Reconfigure: 1hr |
| Observing/Weather Conditions required: | N/A – dome closed. |
| Estimate of effort required | Systems Engineer Mechanical Technician Telescope/Enclosure Operator |

Test Equipment Required: Laser pointer (& gaffer tape) Laser-safe goggles. SAFETY WARNING: observe precautions when handling/observing the laser; to minimise time in which laser is used, can perform in parallel with 4.57 (with 2 laser pointers) Test Procedure or Procedure reference: 1. Affix a laser pointer to Telescope structure near lower edge of Science beam. 2. Move the telescope up and down in elevation and observe position of spot on Windscreen. Success Criterion: Upper edge of windscreen tracks lower edge of science beam. Author : AJB

## ------TVP / 5.6.5.7(1)------ Moonscreen – correct following of Telescope motion

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Moonscreen, ECS. |
| Brief Title of Test | Moonscreen – correct following of Telescope motion |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | Testing of manufacture, operation and range of motion is covered under Enclosure acceptance testing. |
| Estimate of test duration | (During integration) Setup: 1hr (daytime) Test: 1hr (daytime, OSD closed) Analysis time: n/a Reconfigure: 1hr |
| Observing/Weather Conditions required: | N/A – dome closed. |
| Estimate of effort required | Systems Engineer Mechanical Technician Telescope/Enclosure Operator |

Test Equipment Required: Laser pointer (& gaffer tape) Laser-safe goggles. SAFETY WARNING: observe precautions when handling/observing the laser; to minimise time in which laser is used, can perform in parallel with 4.56 (with 2 laser pointers). Test Procedure or Procedure reference: 1. Affix a laser pointer to Telescope structure near upper edge of Science beam. 2. Move the telescope up and down in elevation and observe position of spot on Moonscreen. Success criterion: Edge of moonscreen tracks upper edge of science beam within prescribed range of movement. Author : AJB

## ------TVP / 5.6.5.8(1)------ Flat-field lighting – control from VLT Control Room

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Flat-field lighting |
| Brief Title of Test | Flat-field lighting – control from VLT Control Room |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities | Design/manufacture/test of flat-field lighting is covered under Enclosure Technical Acceptance. |
| Estimate of test duration | (During integration) Setup: n/a Test: 0.5hr (daytime) Analysis time: n/a Reconfigure: n/a |
| Observing/Weather Conditions required: | N/A – dome closed. |
| Estimate of effort required | Systems Engineer VISTA Operator (in VLT Control Room) |

Test Equipment Required: None. Test Procedure or Procedure reference: - Position Operator in the VLT Control Room - Position an Observer on the VISTA Enclosure azimuth floor, in phone/radio contact with Operator. - Select each set of illumination in turn from VLT Control Room (including “off”) - Ensure all states are illuminated as required. - Ensure illumination states and reported correctly to VISTA Operator. Author : AJB all three dimmers (DM43.1 to DM43.3, at the bottom-far-right of cabinet E.H04 on the Az floor), to "Minimum" as abaseline (NOTE: This is a functionality test of lamp control. There is the related issue of whether or not the level of illumination is “suitable” with respect to IR Camera saturation levels etc. (at time of writing, each FF lamp unit is 500W which is proving much too bright; possible reduce to 150W or lower?). System is “compliant” in the sense that it functions in the manner in which it was specified to EIE. Flat-field illumination level will be resolved during IR Camera integration.)

## ------TVP / 5.6.5.8(3)----Flat Field Illumination (Repeatable)

| VISTA System |  |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Flat field illumination – stability. |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | Setup: 10 mins. Test: 1 hr per day, every day for 1 week. Analysis time: 0.5hrs per day, every day for 1 week. Reconfigure: None. |
| Observing/Weather Conditions required: | None – daytime. |
| Estimate of manpower required | 1 operator. 1 VDFS software. |

Test Equipment Required: NOTE: Refer to procedure for 5.6.5.8(1) for details of modified bulb installation and level adjustment. Test Procedure or Procedure reference: 1. Switch on flat-field lamps. 2. Allow 5mins for lamps to stabilise. 3. Repeat: ten consecutive 5sec exposures followed by a 5min wait with lamps on, for a period of one hour, 4. Switch lamps off. 5. Repeat (1)-(4) above every day for 1 week. Reduction/Analysis Procedure: Calculate median counts. Success criteria: Median counts are stable to 0.5% during each single hour. Author : WJS

## ------TVP / 5.6.5.9(2)------ Dome Crane – Can Reach M1 Handling Location

COMPLETED

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Dome Crane |
| Brief Title of Test | Crane can reach handling location. |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: VIS-PRO-EIE-10100-0280 / ??.bmp (check at UKATC)

## ------TVP / 5.6.5.9(11)------ Dome Crane ‘Out of Park’ switch

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Dome |
| Brief Title of Test | Dome Crane ‘Out of Park’ switch |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 0.5 hr Analysis: 0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Functioning dome Crane, Dome System Workstation Test Procedure or Procedure reference: Move the crane out of the ‘Park’ position and verify that the ‘Out of Park’ switch status change is reported to the LCU and the Axes are interlocked out. Author : BS

## ------TVP / 5.6.5.9(12)------ Dome Crane inhibit

| VISTA System | Enclosure |
| --- | --- |
| VISTA Sub-System | Dome |
| Brief Title of Test | Dome Crane inhibit |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | (During integration) Setup: 0 hr Test: 0.5 hr Analysis: 0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Functioning dome control system, external switching device Test Procedure or Procedure reference: Activate the crane external hardware input signal and verify that the crane is inhibited. Author : BS

## ------TVP / 6.1.1------ Control of Equipment from Control Building

| VISTA System | Software |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Control of Equipment from Control Building |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | None additional to other tests |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | N/A |

Test Equipment Required: None Test Procedure or Procedure reference: Perform all procedures that require observations and equipment to be controlled computer equipment in the Paranal Control Room. SPECIFY… Success Criteria: All observing and calibration procedures can performed successfully without requiring intervention in the VISTA Enclosure. Author : JMS

## ------TVP / 6.1.6(2)------ Missing Data or Poor Quality Indicated in the Headers

| VISTA System | Data Handling |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Missing Data or Poor Quality Indicated in the Headers |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 2 hour (During integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software |

Test Equipment Required: None Test Procedure or Procedure reference: Acquire Data files with representative faults, as follows: (a) Telescope vignetted (b) Guide loop closed on non-existent star (c) Telescope not tracking (d) Simulated - science detector disconnected (e) ... plus others TBD Reduction/Analysis Procedure: Check data header items for fault conditions as above. Success Criteria: Each fault is identifiable from the header items. Author : JMS

## ------TVP / 6.1.6(3)------ All Faults Shall Be Logged with Relevant Details

| VISTA System | Software |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | All Faults Shall Be Logged with Relevant Details |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 day (During integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Operator 1 Software |

Test Equipment Required: None Test Procedure or Procedure reference: Induce a representative set of faults artificially, as follows: (a) Trip the M1 maintenance restraint whilst force balancing (b) Depower the M2 Unit electronics and demand an M2 update (c) Trip a telescope interlock whilst tracking (d) Trip an enclosure interlock whilst tracking (e) Disconnect a filter driver cable and demand a filter (f) Disconnect an autoguider whilst the loop is closed (g) Disconnect a LOWFS whilst the loop is closed. Reduction/Analysis Procedure: Examine the fault log. Success Criteria: Each fault is logged with appropriate data available in the same or other log. Author : JMS

## ------TVP / 6.1.8------ Noise interference from telescope mechanisms

| VISTA System | Telescope |
| --- | --- |
| VISTA Sub-System | Camera |
| Brief Title of Test | Noise interference from telescope mechanisms |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | (During integration) Setup: 0.5 hr Test: 0.5 hr Analysis: 3.0 hr Reconfigure: 0 hr |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 1 Systems 1 Operator |

Test Equipment Required: Functioning Camera and mechanisms. Test Procedure or Procedure reference: One at a time, continuously operate those telescope mechanisms that normally function during observations using the local control panels during an observation and examine image data for additional noise compared to operation with no mechanisms operating. Author : BS

## ------TVP / 7.3.5.3------Camera Installation and Removal

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Camera Installation and Removal |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 1 week |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | Systems+Engineer+Safety+Crane Operator |

Test Equipment Required: IR Camera, telescope, enclosure, visible camera. Test Procedure or Procedure reference: Installation and removal procedures for the camera are documented in VIS-PRO-RAL-06067-2002. Full installation requires a total of 5 people. Astrometric stability: Once on-sky operations are resumed and the telescope pointing has been re-established following initialisation of M1, the astrometric stability can be verified as follows: - Take a single image on-sky - Note coordinates of a small number of stars in the corners and centre of the field, and compare these positions to 2MASS coordinates. This can all be done using the RTD image display. Note: For astrometric stability, the specification implies that a series of images should be taken over a period of two nights following cool-down to verify that there are no changes in the pincushion distortion pattern after 48 hours. Reduction/Analysis Procedure: Use the data flow pipeline to confirm the distortion is within limits. Author : GD

## ------TVP / 7.3.5.4------ Camera Intervention

| VISTA System | IR Camera |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test | Camera Intervention |
| Location of Test | Site |
| Applicable Documents: | VIS-SPE-ATC-06000-0004 |
| Preceding Verification Activities |  |
| Estimate of test duration | 10 days (verified during integration) |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required | 3 technicians + Systems |

Test Equipment Required: Telescope , IR Camera, Data analysis workstation Test Procedure or Procedure reference: Procedures for removal and installation of the camera are documented in VIS-PRO-RAL-06067-2002 - Start stopwatch - Commence a warm-up of the camera (needs to be taken off the telescope and into the IPL for this). - Wait until it is safe to open up (e.g. to change a filter) - Start cooling again using LN2 system in IPL. - Stop the stopwatch at the point where the camera is at the temperature range where it would normally be replaced in the telescope. NoteL This takes less than 1 hour once the camera is in the IPL at ambient temperature and pressure. (Tested during June 2007). The filter exchange procedure is documented in VIS-PRO-RAL-06053-9001. Author : GD

## ------TVP / 7.4.12------ Inter-Subsystem Interlock Test

| VISTA System | Safety Interlock System |
| --- | --- |
| VISTA Sub-System | Inter-subsystem interlocks |
| Brief Title of Test | Inter-Subsystem Interlock Test |
| Location of Test | Site |
| Applicable Documents: | AD12 (VLT Electronic Design Specification) |
| Preceding Verification Activities | Verification of interlock operation/reporting within each subsystem is covered within that subsystem’s Technical Acceptance process |
| Estimate of test duration | (During integration) Setup: 2hr (daytime) Test: 1d Analysis time: n/a Reconfigure: 1hr |
| Observing/Weather Conditions required: | N/A – dome closed. |
| Estimate of effort required | Systems Engineer Workpackage Manager - Enclosure Workpackage Manager - Telescope Mechanical Technician Electrical Technician |

Test Equipment Required: None. Test Procedure or Procedure reference: Test operation of interlocks, key overrides as per AD12 above: 1. Telescope↔Crane 2. Enclosure↔Crane 3. Telescope↔Camera Service Trolley … etc. TBD. Reduction/Analysis Procedure: Ensure interlock fault statuses reported correctly. Author : AJB

## ------TVP / Supplementary (1)------ Black Paint Adherence

| VISTA System |  |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test |  |
| Location of Test | Review/Sub-Contractor/Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: Finish Specification VIS-SPE-VER-01001-0440 Rev B Sec. 3.1.2, Adherence test. Test not advised with M1 in place, as test is partially destructive. Test was originally carried out by VRSI subcontractors (ref MIP reports). Author : PFJ

## ------TVP / Supplementary (2)------ LOWFS Image Quality Check

| VISTA System |  |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test |  |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration |  |
| Observing/Weather Conditions required: |  |
| Estimate of manpower required |  |

Test Equipment Required: Test Procedure or Procedure reference: This item essentially identical to “Test 1 – aberration/measurement consistency” of 4.48, (TVP ref: 5.1.18.2), Low Order Wavefront Sensor. Author: WJS.

## ------TVP / Supplementary (3)------ System Image Quality Check

| VISTA System |  |
| --- | --- |
| VISTA Sub-System |  |
| Brief Title of Test |  |
| Location of Test | Site |
| Applicable Documents: |  |
| Preceding Verification Activities |  |
| Estimate of test duration | 3 nights + 3 days analysis. |
| Observing/Weather Conditions required: | Good seeing. |
| Estimate of manpower required | Operator. Scientist. WFS software expert. |

Test Equipment Required: Test Procedure or Procedure reference: TEST 1: Absolute zeropoints of the HOWFS and LOWFS : 1. Using the IR Camera, compare LOWFS and HOWFS measurements taken quasi-simultaneously. 2. the null off-axis aberrations from differing off-axis position are automatically subtracted, so these should be consistent. TEST 2: While rolling the IR Camera in 10 degree steps from 0,10,20,...,180 deg, perform “Test 1 – aberration/measurement consistency” of 4.48, (TVP ref: 5.1.18.2), Low Order Wavefront Sensor. Reduction/Analysis Procedure: 1. Confirm aberrations introduced match those measured. Notes: - Telescope aberrations should be nearly constant during the track , so when the LOWFS and HOWFS Zernike measurements are converted back to the M1 coordinate system they should be independent of Camera roll angle. - Sinusoidal residuals may indicate a bias in the LOWFS/HOWFS , or non-nominal aberrations internal to the Camera lens barrel: if so, further mapping using 8 different HOWFS positions across the field can help to diagnose this. Author : WJS

__oOo__